## Slide 1: Operation of the Load vs. Displacement Analyzer
Joshua McGuckin
Associate Research Engineer

## Slide 2: Range of Motion (ROM)
The 6-degrees of freedom motion simulator (6DOF) allows us to quantify total range of motion (ROM) in the 6 directions
    Flexion-Extension (FE)
    Lateral Bending (LB)
    Axial Rotation (AR)
Tells us how much spinal segments move during applied bending
Load cell cranial to the specimen measures applied torque in the direction of bending
Pneumatics ensure weightlessness of the specimen to achieve pure bending moments
  i.e. equal torque at every level of the spine
ML Translation
AP Translation
CC translation
Lateral Bending
Axial Rotation
Flexion Extension

## Slide 3: Range of Motion (ROM)
The 6-degrees of freedom motion simulator (6DOF) allows us to quantify total range of motion (ROM) in the 6 directions
    Flexion-Extension (FE)
    Lateral Bending (LB)
    Axial Rotation (AR)
Tells us how much spinal segments move during applied bending
Load cell cranial to the specimen measures applied torque in the direction of bending
Pneumatics ensure weightlessness of the specimen to achieve pure bending moments
  i.e. equal torque at every level of the spine
  This is not yet validated!!!
ML Translation
AP Translation
CC translation
Lateral Bending
Axial Rotation
Flexion Extension

## Slide 4: Range of Motion (ROM)
The 6-degrees of freedom motion simulator (6DOF) allows us to quantify total range of motion (ROM) in the 6 directions
    Flexion-Extension (FE)
    Lateral Bending (LB)
    Axial Rotation (AR)
Tells us how much spinal segments move during applied bending
ROM

## Slide 5: Range of Motion (ROM)
The 6-degrees of freedom motion simulator (6DOF) allows us to quantify total range of motion (ROM) in the 6 directions
    Flexion-Extension (FE)
    Lateral Bending (LB)
    Axial Rotation (AR)
Tells us how much spinal segments move in each direction during applied bending
But there’s so much MORE!
ROM

## Slide 6: BEYOND ROM
ROM is moving in response to Torque!!!
  We can think of torque as load and ROM as deformation
Spines aren’t linear elastic (at least, not always)
  Instead, we see an S-shaped hysteresis curve
  Because this curve is non-linear, ROM is not the same stiffness (slope of this curve)
  Specimens have different stiffness at different parts of the curve
  In fact, specimens with the same ROM can have different stiffnesses and vice-versa
  Why is this important?
    ROM is a limited parameter that can’t answer all questions about spinal biomechanics
    We often use ROM as a metric to elude to other things like segmental stiffness, instability, motion preservation
    But this is not always a direct measurement
  With hysteresis curves, we can directly measure:
    Instability, stiffness, non-linearity, energy dissipation

## Slide 7: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?

## Slide 8: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
How do we analyze it???

## Slide 9: So you want to Analyze S-curves…
First need to know what an S-curve is...
    Mechanical properties are time-, load-, and deformation-dependent
    Also history-dependent, which is why the 6DOF preconditions the specimen with 2 cycles in FE, LB, AR
    For our purposes, only data from the 3rd cycle is analyzed
  In spine biomechanics, ideal hysteresis curves resemble S-shapes1
    Plot Load (Torque) vs. Deformation (Relative Angle)
    Biphasic behavior (bottom-right and top-left curves)
      Represents two-directional bending/torsion in the plane of interest
        Flexion-Extension
        Right LB-Left LB
        Left AR-Right AR

## Slide 10: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ)1:
      A measure of laxity of the spinal specimen
      Describes the range over which the spine moves freely without load
        i.e. under its own weight
        Used to measure instability compared to intact
          LARGER NZ = MORE instability
      Defined as the difference in angulation at zero-load between the 2-phases
        These 2-points represent the positive (top) and negative (bottom) Neutral Zone Boundaries (NZBs)

## Slide 11: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ)1:
      Deformation outside the NZ
        From the end of the NZ to the maximum load and minimum load
          Can think of these as positive (top) and negative (bottom) Ezs
      There is no general definition to the EZ
        It is generally thought of as the region outside of the NZ
        But by the definition as the “stiffest region”, the analyzer application described here segments the region between the NZB and the min/max ROM into 3 parts based on their stiffnesses.
        Here, the region with the greatest stiffness is formally defined as EZ

## Slide 12: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ):
    Neutral Position (NP)1:
      NP is the midpoint between these 2 points of angulation at zero-load
        i.e. midpoint of NZ
        Will ideally be close to the origin (0,0)
          Unless drift or misaligned Load and ROM data

## Slide 13: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ):
    Neutral Position (NP):
    Range of Motion (ROM)1:
      Total angulation of the specimen in the plane of motion (FE, LB, AR)
      Describes how much motion there is in a specimen/construct
      Can be split into positive & negative ROM based on the NP

## Slide 14: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ):
    Neutral Position (NP):
    Range of Motion (ROM):
    Neutral Zone Stiffness (NZS)1:
      Represents the lax deformation of the specimen/construct
        Lowest stiffness in the S-curve
      Is biphasic, just like NZ, so there is a positive (right) and a negative (left) NZS
      Defined as the slope (deformation/load)of the NZ

## Slide 15: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ):
    Neutral Position (NP):
    Range of Motion (ROM):
    Neutral Zone Stiffness (NZS):
    Elastic Zone Stiffness (EZS)1:
      Represents the elastic deformation of the specimen/construct
      Is biphasic, just like EZ, so there is a positive (top) and negative (bottom) EZS
      Defined as the slope at the stiffest region of the EZ starting from the max/min load

## Slide 16: So you want to Analyze S-curves…
  Important Terms
    Neutral Zone (NZ):
    Elastic Zone (EZ):
    Neutral Position (NP):
    Range of Motion (ROM):
    Neutral Zone Stiffness (NZS):
    Elastic Zone Stiffness (EZS):
    Energy Dissipation1:
      Represents the viscoelastic/plastic behavior of the specimen/construct
      Defined by the area under the hysteresis curve (can be normalized to ROM)2,5
        Greater area = more viscous behavior
        Less area = more elastic behavior

## Slide 17: So you want to Analyze S-curves…
  Other Terms
    Sigmoidity (NZS/EZS)
      A measure of non-linearity of the mechanical properties of the specimen or construct
        How much stiffness changes in the specimen
        Greater sigmoidity = more linear elastic
        Less sigmoidity = more non-linear

## Slide 18: So you want to Analyze S-curves…
  Other Terms
    Instability
      Often when the specimen is injured/degenerated
      Considered unstable when ROM or NZ significantly changes from intact1,3
        NZ is considered a more direct estimate of instability4
          NZ is considered more sensitive to detecting changes in motion due to spondylolisthesis and fracture4,5
          BUT findings with NZ depend on which method is used to find it, meaning if you find NZ with one method, you may not notice the same values or findings as if you were to analyze it with another method5
          This analyzer used the Zero-Load method described by Wilke et al1
          Additional methods include:
              Double Sigmoid
              Stiffness Threshold
              Extrapolated Elastic Zone
      NZ Ratio or NZR (NZ/ROM)2
        Used when ROM differs between specimens and groups
        Describes how much of the motion is “unstable”

## Slide 19: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?

## Slide 20: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?

## Slide 21: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
  With hysteresis curve analysis we know:
    NZ increased  MORE instability

## Slide 22: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
  With hysteresis curve analysis we know:
    NZ increased  MORE instability
    EZS increased  stiffer at extreme ROM, likely due to stretching of the disc and ligaments

## Slide 23: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
  With hysteresis curve analysis we know:
    NZ increased  MORE instability
    EZS increased  stiffer at extreme ROM, likely due to stretching of the disc and ligaments
    Energy dissipation increased  more load focused on the disc

## Slide 24: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
  With hysteresis curve analysis we know:
    NZ increased  MORE instability
    EZS increased  stiffer at extreme ROM, likely due to stretching of the disc and ligaments
    Energy dissipation increased  more load focused on the disc
    Sigmoidity increased  more non-linear, more viscous behavior, not-as-smooth motion (more abrupt transitions between NZ and EZ)

## Slide 25: BEYOND ROM: Example
This is data from a T12-L1 specimen
There is similar ROM, but clearly, these curves are different
Can we still say motion is the same?
  With hysteresis curve analysis we know:
    NZ increased  MORE instability
    EZS increased  stiffer at extreme ROM, likely due to stretching of the disc and ligaments
    Energy dissipation increased  more load focused on the disc
    Sigmoidity increased  more non-linear, more viscous behavior, not-as-smooth motion (more abrupt transitions between NZ and EZ)
      Ex: In a fused segment, expect lower instability, lower sigmoidity, low energy dissipation

## Slide 26: Value of Hysteresis Curves
More detail than ROM, ROM does not represent all of spine biomechanics
More sensitive to changes in stability, stiffness, and certain anatomical changes
Can tell you things beyond total motion
    How stiff is the construct? (NZS, EZS)
    How stable is the construct? (NZ)
    How smooth is the motion? (sigmoidity)
    What is being loaded? Soft tissue or bony tissue? (sigmoidity, energy dissipation)
Sometimes ROM is enough
    Pre- vs. Post-fatigue
    Intact vs. Construct
Surgeons, PD, and Marketing won’t know what any of these metrics mean, other than ROM, so we have to explain it to them very clearly if we use them
But if what you are trying to capture needs a more sensitive metric or you want to explore the data more in-depth, use this analyzer

## Slide 27: Value of Hysteresis Curves
Pretty pictures for your publications
More data to choose from to support your conclusions
With more data, we can say MORE about our products
  Improves stiffness
  Increases stability
  Makes transition of motion smoother
  Conducts loads more efficiently
Makes for a much more compelling paper

## Slide 28: Limitations
Several assumptions are made that may not be completely accurate under all testing conditions
  Pure bending moments:
    Equal torque is applied to all levels, no resultant compression or tension, no axial compression
    Allows us to use the SAME Load data at all levels and to analyze hysteresis curves for single-level and whole-spine OptoTrak data
    Can analyze single- and multi-level constructs, if true, but focus your analysis on the construct levels
  Hysteresis curve is in a normal S-shape
    1st phase is below the 2nd phase, posNZB is always above the negNZB
    No abnormal noise changing the curve shape
  There are only 2 Zero-Load points
    The spine only returns to a weightless position 2X

## Slide 29: Limitations
If any of these assumptions are not met, this application is far less accurate
  Ex. If pure bending moments are not applied, the torque data is not universal to every level
Additionally limited by alignment of OptoTrak & Load datasets
  Improper alignment will result in incorrect hysteresis curves and inaccurate data
  It is the responsibility of the user to validate
Also limited in ROM accuracy
  The application smoothens and reduces the resolution of OptoTrak data
  As such, ROM is typically reported as 0-2⁰ less than it actually is, so use the 6DOM ROM ANALYZER to accurately report ROM

## Slide 30: References
  Wilke HJ, Wenger K, Claes L. Testing criteria for spinal implants: recommendations for the standardization of in vitro stability testing of spinal implants. Eur Spine J. 1998;7(2):148-154. doi:10.1007/s005860050045
  Zirbel SA, Stolworthy DK, Howell LL, Bowden AE. Intervertebral disc degeneration alters lumbar spine segmental stiffness in all modes of loading under a compressive follower load. The Spine Journal. 2013;13(9):1134-1147. doi:10.1016/j.spinee.2013.02.010
  Wilke HJ, Herkommer A, Werner K, Liebsch C. In vitro analysis of the segmental flexibility of the thoracic spine. PLoS One. 2017;12(5):e0177823. Published 2017 May 16. doi:10.1371/journal.pone.0177823
  Tan E, Wang T, Pelletier MH, Walsh WR. Effects of cement augmentation on the mechanical stability of multilevel spine after vertebral compression fracture. Journal of Spine Surgery. 2016;2(2):111-121. doi:10.21037/jss.2016.06.05
  Di Pauli von Treuheim T, Zengerle L, Hecht AC, Iatridis JC, Wilke HJ. Does the neutral zone quantification method matter? Efficacy of evaluating neutral zone during destabilization and restabilization in human spine implant testing. J Biomech. 2021;129:110756. doi:10.1016/j.jbiomech.2021.110756

## Slide 31: Installing Load vs. Displacement analyzer
Install MATLAB Runtime R2022a (9.12)
  Need this version to run the application

## Slide 32: Running Load vs. Displacement analyzer
Go to the server and start the application
  \\SW10FIL01\Research\Research Operations\Biomechanical\Software\LoadvsDisp Analysis Tools\Version 2.4\Load_vs_Displacement_Analyzer\for_testing

## Slide 33: Running Load vs. Displacement analyzer
OR use the MATLAB code
  \\SW10FIL01\Research\Research Operations\Biomechanical\Software\LoadvsDisp Analysis Tools\Version 2.4 - Current App Version.load_disp_edits_v2_4.m
  Just make sure to go to add-ons and verify that the Statistics and Machine Learning toolbox and Signal Processing toolbox are installed

## Slide 34: Running Load vs. Displacement analyzer
The application will start as indicated by the following popup icon

## Slide 35: Running Load vs. Displacement analyzer
Select the motor data
  Select the folder containing the load files (Books) for the construct being analyzed

## Slide 36: Running Load vs. Displacement analyzer
Select the OptoTrak data
  Select the folder containing the cal.xls files for the construct being analyzed

## Slide 37: Running Load vs. Displacement analyzer
Name the construct groups
Fill in the specimen number and construct with any info needed to identify the construct

## Slide 38: Running Load vs. Displacement analyzer
Select the files for analysis
Select any files that are of interest in your analysis
  These books correspond to FE, LB, AR for the specimens in the selected folder in order (always in this order)
  It is recommended to only select one at a time and then re-run the application for the remaining files to avoid not having your work saved if there is a problem with one of these files, the program only outputs a data summary after it has successfully analyzed all the selected files

## Slide 39: Running Load vs. Displacement analyzer
Progress bar
A progress bar will pop up indicating where the program is with the analysis

## Slide 40: Running Load vs. Displacement analyzer
Verify the alignment of the data
A new popup figure will show the hysteresis curve (top) and the 3rd cycle of OptoTrak vs. Time and Load vs. Time data (bottom) as seen by the application
Make sure that:
  The hysteresis curve resembles the distinctive S-shape
  The peaks from the OptoTrak data align with those of the Load data
Failure to ensure 2 will result in an inaccurate hysteresis curve

## Slide 41: Running Load vs. Displacement analyzer
Verify the alignment of the data
This step is crucial and can be prone to error
  To identify the 3rd cycles, the application counts the number of peaks and valleys and determines the starting point for each cycle as when the OptoTrak vs. Time or Load vs. Time data intersects zero (i.e. the x-axis) which is representative of the starting position or load, respectively
  Because the OptoTrak Motion Capture System collects data at a much higher rate (~100 Hz) than the 6DOF Load Cell (~10 Hz), much of the OptoTrak data needs to be negated in order for these two datasets to be aligned
  Additionally, the OptoTrak Motion Capture System starts collecting before the 6DOF Load Cell which causes a delay in the OptoTrak vs. Time data that the application needs to “chop off” in order to align the two datasets
  There can calls be extra data at the end of the OptoTrak vs. Time dataset because this is turned off after the 6DOF Load Cell terminates data collection

## Slide 42: Running Load vs. Displacement analyzer
Verify the alignment of the data
This step is crucial and can be prone to error
  The algorithm does its best to chop off this excess OptoTrak data on its own, but sometimes it is not enough and the user needs to manually edit the data in Excel
  The failure of the algorithm to chop off the data causes it to misinterpret which cycle is the last cycle and where said cycle starts, which affects the hysteresis curve

## Slide 43: Running Load vs. Displacement analyzer
Verify the alignment of the data
This step is crucial and can be prone to error
  The algorithm does its best to chop off this excess OptoTrak data on its own, but sometimes it is not enough and the user needs to manually edit the data in Excel
  The failure of the algorithm to chop off the data causes it to misinterpret which cycle is the last cycle and where said cycle starts, which affects the hysteresis curve

## Slide 44: Running Load vs. Displacement analyzer
Verify the alignment of the data
This step is crucial and can be prone to error
  The algorithm does its best to chop off this excess OptoTrak data on its own, but sometimes it is not enough and the user needs to manually edit the data in Excel
  The failure of the algorithm to chop off the data causes it to misinterpret which cycle is the last cycle and where said cycle starts, which affects the hysteresis curve

## Slide 45: Running Load vs. Displacement analyzer
Manual Data Selection
Press “Enter” and a new popup message will appear
Press “YES” if the hysteresis curve looks off or there is misalignment between the peaks of the OptoTrak vs. Time and Load vs Time data
Otherwise press “NO” and continue to slide 49

## Slide 46: Running Load vs. Displacement analyzer
Manual Data Selection: “YES”
If “YES” is pressed, a new popup figure will appear showing the OptoTrak vs. Time data with the reduced resolution (~10 Hz)
Follow the instructions to manually select the start and end points of the 3rd cycle of OptoTrak vs. Time data

## Slide 47: Running Load vs. Displacement analyzer
Manual Data Selection: “YES”
A new popup figure will appear showing the Load vs. Time data
Follow the instructions to manually select the start of the 3rd cycle
  As two vectors that will be combined, the OptoTrak vs. Time and Load vs. Time datasets need to have the same number of data points
  The endpoint of the 3rd cycle of the Load vs. Time dataset will be determined by the number of data points between the selected start and end points in the 3rd cycle of the OptoTrak vs. Time dataset

## Slide 48: Running Load vs. Displacement analyzer
Manual Data Selection: “YES”
A new popup figure will appear showing the updated hysteresis curve
Load vs. Time data will remain the same, but raw OptoTrak vs. Time data will be plotted with the lag
  Shows the data as it was pre-analysis
The 3rd cycle of the OptoTrak vs. Time data will be shifted based on the starting point selected in the previous two figures and should be aligned with the 3rd cycle of the Load vs. Time data
Press “YES” if the data alignment and the hysteresis curve are satisfactory (continue to slide 49)
  Note: This type of “lag chopping” can be done outside of the application in the cal files
Otherwise, press “NO” and re-select the 3rd cycle of OptoTrak vs. Time and Load vs. Time (slides 46-47)

## Slide 49: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection:
If the 3rd cycle of OptoTrak vs. Time and Load vs. Time data are aligned and properly selected, yet the hysteresis curve still looks odd press “YES”
  Reasons for abnormal hysteresis curves:
    Discontinuities: Common at very low ROM and drift
    Noise: This application finds the NZ using the Zero Load method described by Wilke et al1 but there are many others5 that could be used. Therefore, if there are more than two intersections with the y-axis, often due to noisy data, the application cannot identify the NZ
    Non-sensical data: segment anatomy or loading rate prevents the smooth transition between cycles, resulting in a hysteresis curve that doesn’t resemble an S-shape
Otherwise, press “NO” and continue to slide 66
  The hysteresis curve looks as it should or as good as it will ever look

## Slide 50: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection: “YES”
A new popup figure will appear showing the hysteresis curve
  The dashed line represents the x-axis that is used in the Zero Load method
  Ideally the hysteresis curve starts and ends around the origin (0,0)
    Seen in the 1st phase (bottom-right) of the hysteresis curve as the 2 points on opposite sides of the x-axis, closest to the x-axis

## Slide 51: Running Load vs. Displacement analyzer
3rd Cycle Start
3rd Cycle End

## Slide 52: Running Load vs. Displacement analyzer
3rd Cycle Start
3rd Cycle End
1st Phase
2nd Phase

## Slide 53: Running Load vs. Displacement analyzer
3rd Cycle Start
3rd Cycle End

## Slide 54: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection: “YES”
Zoom in, press “Enter”, and select the positive NZB (posNZB)
  A new dashed line will appear at the posNZB
  Remember, this is where the 2nd phase (top-left) of the hysteresis curve intersects with the x-axis (i.e. zero-load)

## Slide 55: Running Load vs. Displacement analyzer

## Slide 56: Running Load vs. Displacement analyzer
X-axis
Zero-Load
2nd phase point closest to X-axis
posNZB

## Slide 57: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection: “YES”
Press “Enter” and select the negative NZB (negNZB)
  A new dashed line will appear at the negNZB
  If there is a discontinuity, select the starting point of the hysteresis curve as negNZB

## Slide 58: Running Load vs. Displacement analyzer

## Slide 59: Running Load vs. Displacement analyzer
X-axis
Zero-Load
1st phase point closest to X-axis
(always the origin)
posNZB
negNZB

## Slide 60: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection: “YES”
Press “Enter” and select the points of the 1st phase that intersect the posNZB
  All points between the posNZB and negNZB in the first phase will turn red

## Slide 61: Running Load vs. Displacement analyzer

## Slide 62: Running Load vs. Displacement analyzer
posNZB
negNZB
1st phase
Points where 1st phase intersects posNZB

## Slide 63: Running Load vs. Displacement analyzer
Manual Neutral Zone Selection: “YES”
Press “Enter” and select the points of the 2nd phase that intersect the posNZB
  All points between the posNZB and negNZB in the 2nd phase will turn red before the figure closes itself

## Slide 64: Running Load vs. Displacement analyzer

## Slide 65: Running Load vs. Displacement analyzer
posNZB
negNZB
2nd Phase
Points where 2nd phase intersects negNZB

## Slide 66: Running Load vs. Displacement analyzer
Analyzed Hysteresis Curve
The hysteresis curve will be annotated with the features assessed during the analysis:
      posNZB: Top “*”
      posNZB: Bottom “*”
      NZ: red region
      posNZS: right-most black line
      negNZS: left-most black line
      posEZ: top magenta region
      negEZ: bottom magenta region
      posEZS: top green line
      posEZS: bottom green line
      Hysteresis: yellow area within the curve
This figure is saved automatically for visual confirmation of the extracted features

## Slide 67: Running Load vs. Displacement analyzer
Run Again or Close
After the application has analyzed all the selected OptoTrak and Load files, a final pop-up will appear asking the user if they would like to run the script again.
Press “Restart” to re-analyze or analyze more files
Press “Done” to close the application
Before choosing, make sure to verify the accuracy of the output data

## Slide 68: Running Load vs. Displacement analyzer
Verify Data
In the folder containing the OptoTrak data used in this analysis, there should appear a new Excel spreadsheet called “Specimen Summary – Hysteresis Curve.xls”

## Slide 69: Running Load vs. Displacement analyzer
Verify Data
In this spreadsheet, there will be a tab for each relative analyzed
In each tab there will be a table with rows for every cal and Book file combination analyzed (FE, LB, or AR for every construct) with all the extracted data from the hysteresis curve
From these parameters other metrics can be calculated
    Sigmoidity: NZS/EZS
    Normalized NZ: NZ/ROM
It should be noted that because the resolution of the OptoTrak data is reduced to make it compatible with the Load data, the ROM in the hysteresis curve might be less than the ROM of the raw OptoTrak data
Differences between raw and process ROM are typically between 0-2⁰ and must be evaluated before using ROM in any part of the greater analysis

## Slide 70: Running Load vs. Displacement analyzer
Verify Data
In the folder containing the Load data used in this analysis, there should appear a new folder called “Angular Displacement VS Load Plots”
In this folder, there will be 3 types of files for every cal and Book file combination analyzed
These will include JPEGs, MATLAB Figs, and Excel spreadsheets
  JPEGs and MATLAB Figs have the annotated hysteresis curves for reference and verification of data

## Slide 71: Running Load vs. Displacement analyzer
Verify Data
In the folder containing the Load data used in this analysis, there should appear a new folder called “Angular Displacement VS Load Plots”
In this folder, there will be 3 types of files for every cal and Book file combination analyzed
These will include JPEGs, MATLAB Figs, and Excel spreadsheets
  JPEGs and MATLAB Figs have the annotated hysteresis curves for reference and verification of data
  Excel spreadsheets have all the processed data (3 cycles of time, displacement, load) and the data for the hysteresis curve (last cycle time, displacement, and load) that was analyzed by the application
    These can be used to re-create and manually re-analyze the hysteresis curves if necessary
    If the hysteresis curve looks as expected, the user can usually trust the results; however, if the curve looks “wonky”, the user should consider correcting the OptoTrak data in the application (slides 46-47) or in the cal or Book files or manually extracting parameters in the output Excel spreadsheets

## Slide 72: Running Load vs. Displacement analyzer
Fatal Errors
If, for whatever reason, the application encounters a fatal error that prevents it from running the analysis, a warning message will pop up before the program closes
There are a number of reasons that this may occur, but some of the most common include:
  Problems accessing data on the server, often because path names are too long
  Discontinuities in the OptoTrak data usually due to blocked rigid bodies
  Severely abnormal OptoTrack data where peaks cannot be determined (e.g. high drift or extremely low ROM)
  Extra columns in the OptoTrak data that are not part of the original output (e.g. when normalizing relatives)
  Mismatch in the number of OptoTrak and Load files
  Too much lag in the OptoTrak data